## Summary
| Pipette | H2O |  | Blood |  | Methanol |  |  | Code | Description | Volume range | Order Number
|  | µl | r-% | µl | r-% | µl | r-% | Pipette | AA100 | Andrew Alliance Picus 120 µl | 5-120 µl | LH735.041AA | Scale: Mettler Toledo XS205, tared after each measurement
1 | AA120/O/F/H | 100.3 | 0.18257 | 100.4506 | 1.1401 | 97.7089 | 0.55453 |  | AA1000 | Andrew Alliance Picus 1000 µl | 50-1000 µl | LH735.081AA | Evaporation trap applied in all measurements
2 | AA120/O/F/N | 100.2498 | 0.15819 | 98.092 | 0.5689 | 96.6456 | 0.37683 | [5 empty] | Density of matrices determined with pyknometer (20 °C)
3 | AA120/O/R/H | 100.4003 | 0.18239 | 99.8945 | 0.64376 | 98.2405 | 0.46437 | [5 empty] | Room temperature 20 °C
4 | AA120/O/R/N | 100.3201 | 0.1229 | 99.8466 | 0.63728 | 100.0886 | 0.43423 | Tip | O | Optifit | 0,5-200 µl | 790200
5 | AA120/OLR/F/H | 100.4204 | 0.06317 | 102.2052 | 1.4983 | 97.2785 | 0.63229 |  | OLR | Optifit Low Retention | 0,5-200 µl | LH-L790200
6 | AA120/OLR/F/N | 100.0492 | 0.085197 | 99.6261 | 0.99129 | 99.7089 | 0.4973 |  | O | Optifit | 10-1000 µl | 791000
7 | AA120/OLR/R/H | 100.2498 | 0.10807 | 101.7354 | 1.5055 | 97.8861 | 0.64529 |  | OLR | Optifit Low Retention | 10-1000 µl | LH-L791000
8 | AA120/OLR/R/N | 100.3602 | 0.10743 | 99.0604 | 1.33 | 98.4684 | 0.56018 |  | OWB | Optifit Wide Bore | 10-1000 µl | 791020
9 | AA1000/O/F/H | 1007.0621 | 0.07501 | 1012.5216 | 0.3214 | 1002.0886 | 0.31392 | Technique | F | Forward
10 | AA1000/O/F/N | 1003.6923 | 0.038685 | 1010.767 | 0.1275 | 993.481 | 0.32497 |  | R | Reverse
11 | AA1000/O/R/H | 1003.953 | 0.13454 | 1003.72 | 0.29338 | 993.8228 | 0.18708 | Speed | H | Slow (setting: 1)
12 | AA1000/O/R/N | 999.8913 | 0.090456 | 996.117 | 0.46823 | 983.6329 | 0.13388 |  | N | Fast (setting: 9)
13 | AA1000/OLR/F/H | 1005.0161 | 0.053934 | 1003.0105 | 0.23359 | 995.5443 | 0.27077
14 | AA1000/OLR/F/N | 1004.7855 | 0.047005 | 1008.092 | 0.14058 | 996.2278 | 0.20626
15 | AA1000/OLR/R/H | 1004.1436 | 0.092645 | 992.7421 | 0.1469 | 995.3418 | 0.083459 | Blood
16 | AA1000/OLR/R/N | 1000.8741 | 0.076136 | 992.4161 | 0.19543 | 993.2785 | 0.05287 | Tip | O | SafetySpace Filter Tip | 5-200 µl | 790201F
17 | AA1000/OWB/F/H | 1007.7139 | 0.045486 | 998.7536 | 0.20748 | 1001.7975 | 0.18559 |  | OLR | SafetySpace Filter Tip Low Retention | 5-200 µl | LH-LF790201
18 | AA1000/OWB/F/N | 1004.1235 | 0.075156 | 1006.4334 | 1.104 | 988.1139 | 0.50754 |  | O | SafetySpace Filter Tip | 50-1000 µl | 791001F
19 | AA1000/OWB/R/H | 1004.3442 | 0.10326 | 993.9789 | 0.2583 | 993.2405 | 0.11323 |  | OLR | SafetySpace Filter Tip Low Retention | 50-1000 µl | LH-LF791001
20 | AA1000/OWB/R/N | 1000.0417 | 0.045652 | 990.8054 | 0.2931 | 987.8354 | 0.20648 |  | OWB | Optifit Wide Bore, with Safe-Cone in the pipette tip | 10-1000 µl | 791020

## H2O
1 | Andrew Alliance Picus 5-120 µL |  |  |  | 9 | Andrew Alliance Picus 50-1000 µL |  |  |  | Matrix: | Water, deionized
| Sartorius Optifit Tip 0,5-200 µL |  |  |  |  | Sartorius Optifit Tip 10-1000 µL |  |  |  |  | Temperature, Water: 20,5 °C
|  |  | 100 | 100.3% |  |  |  | 1000 | 100.7% |  | Air pressure 1018 hPa
| Average |  |  | 100.3 |  | Average |  |  | 1007.0621 |  | Air moisture: 31 %
| Standard deviation |  |  | 0.18312 |  | Standard deviation |  |  | 0.7554 | Z-factor | 1.003 | mg/µL
|  | [g] | [mg] | [µL] |  |  | [g] | [mg] | [µL]
| 0,0996 99,90 | 0.0996 | 99.6 | 99.8988 |  | 1,0020 1005,01 | 1.002 | 1002 | 1005.006 | Technique: | Forward
| 0,0999 100,20 | 0.0999 | 99.9 | 100.1997 |  | 1,0045 1007,51 | 1.0045 | 1004.5 | 1007.5135 | Speed: | Slow
| 0,0999 100,20 | 0.0999 | 99.9 | 100.1997 |  | 1,0045 1007,51 | 1.0045 | 1004.5 | 1007.5135
| 0,1001 100,40 | 0.1001 | 100.1 | 100.4003 |  | 1,0046 1007,61 | 1.0046 | 1004.6 | 1007.6138
| 0,1000 100,3 | 0.1 | 100 | 100.3 |  | 1,0040 1007,01 | 1.004 | 1004 | 1007.012
| 0,1002 100,50 | 0.1002 | 100.2 | 100.5006 |  | 1,0043 1007,31 | 1.0043 | 1004.3 | 1007.3129
| 0,0999 100,20 | 0.0999 | 99.9 | 100.1997 |  | 1,0043 1007,31 | 1.0043 | 1004.3 | 1007.3129
| 0,1001 100,40 | 0.1001 | 100.1 | 100.4003 |  | 1,0040 1007,01 | 1.004 | 1004 | 1007.012
| 0,1001 100,40 | 0.1001 | 100.1 | 100.4003 |  | 1,0040 1007,01 | 1.004 | 1004 | 1007.012
| 0,1002 100,50 | 0.1002 | 100.2 | 100.5006 |  | 1,0043 1007,31 | 1.0043 | 1004.3 | 1007.3129
| Systematic error |  | µL | 0.3 |  | Systematic error |  | µL | 7.0621
| Systematic error |  | % | 0.2991 |  | Systematic error |  | % | 0.70126
| Random error |  | µL | 0.18312 |  | Random error |  | µL | 0.7554
| Random error |  | % | 0.18257 |  | Random error |  | % | 0.07501
| Uncertainty |  | µL | 0.66624 |  | Uncertainty |  | µL | 8.573
2 | Andrew Alliance Picus 5-120 µL |  |  |  | 10 | Andrew Alliance Picus 50-1000 µL |  |  |  | Matrix: | Water, deionized
| Sartorius Optifit Tip 0,5-200 µL |  |  |  |  | Sartorius Optifit Tip 10-1000 µL |  |  |  |  | Temperature, Water: 20 °C
|  |  | 100 | 100.2% |  |  |  | 1000 | 100.4% |  | Air pressure 1018 hPa
| Average |  |  | 100.2498 |  | Average |  |  | 1003.6923 |  | Air moisture: 31 %
| Standard deviation |  |  | 0.15859 |  | Standard deviation |  |  | 0.38828 | Z-factor | 1.0029 | mg/µL
|  | [g] | [mg] | [µL] |  |  | [g] | [mg] | [µL]
| 0,0998 100,10 | 0.0998 | 99.8 | 100.0994 |  | 1,0008 1003,70 | 1.0008 | 1000.8 | 1003.7023 | Technique: | Forward
| 0,0999 100,20 | 0.0999 | 99.9 | 100.1997 |  | 1,0015 1004,40 | 1.0015 | 1001.5 | 1004.4043 | Speed: | Fast
| 0,0998 100,10 | 0.0998 | 99.8 | 100.0994 |  | 1,0009 1003,80 | 1.0009 | 1000.9 | 1003.8026
| 0,0997 100,00 | 0.0997 | 99.7 | 99.9991 |  | 1,0005 1003,40 | 1.0005 | 1000.5 | 1003.4014
| 0,1000 100,3 | 0.1 | 100 | 100.3 |  | 1,0013 1004,20 | 1.0013 | 1001.3 | 1004.2038
| 0,1000 100,3 | 0.1 | 100 | 100.3 |  | 1,0009 1003,80 | 1.0009 | 1000.9 | 1003.8026
| 0,1001 100,40 | 0.1001 | 100.1 | 100.4003 |  | 1,0006 1003,50 | 1.0006 | 1000.6 | 1003.5017
| 0,1002 100,50 | 0.1002 | 100.2 | 100.5006 |  | 1,0002 1003,10 | 1.0002 | 1000.2 | 1003.1006
| 0,0999 100,20 | 0.0999 | 99.9 | 100.1997 |  | 1,0007 1003,60 | 1.0007 | 1000.7 | 1003.602
| 0,1001 100,40 | 0.1001 | 100.1 | 100.4003 |  | 1,0005 1003,40 | 1.0005 | 1000.5 | 1003.4014
| Systematic error |  | µL | 0.24985 |  | Systematic error |  | µL | 3.6923
| Systematic error |  | % | 0.24923 |  | Systematic error |  | % | 0.36787
| Random error |  | µL | 0.15859 |  | Random error |  | µL | 0.38828
| Random error |  | % | 0.15819 |  | Random error |  | % | 0.038685
| Uncertainty |  | µL | 0.56703 |  | Uncertainty |  | µL | 4.4688
3 | Andrew Alliance Picus 5-120 µL |  |  |  | 11 | Andrew Alliance Picus 50-1000 µL |  |  |  | Matrix: | Water, deionized
| Sartorius Optifit Tip 0,5-200 µL |  |  |  |  | Sartorius Optifit Tip 10-1000 µL |  |  |  |  | Temperature, Water: 20 °C
|  |  | 100 | 100.4% |  |  |  | 1000 | 100.4% |  | Air pressure 1018 hPa
| Average |  |  | 100.4003 |  | Average |  |  | 1003.953 |  | Air moisture: 31 %
| Standard deviation |  |  | 0.18312 |  | Standard deviation |  |  | 1.3507 | Z-factor | 1.0029 | mg/µL
|  | [g] | [mg] | [µL] |  |  | [g] | [mg] | [µL]
| 0,1002 100,50 | 0.1002 | 100.2 | 100.5006 |  | 1,0022 1005,11 | 1.0022 | 1002.2 | 1005.1064 | Technique: | Reverse
| 0,1002 100,50 | 0.1002 | 100.2 | 100.5006 |  | 1,0018 1004,71 | 1.0018 | 1001.8 | 1004.7052 | Speed: | Slow
| 0,1003 100,60 | 0.1003 | 100.3 | 100.6009 |  | 1,0008 1003,70 | 1.0008 | 1000.8 | 1003.7023
| 0,0998 100,10 | 0.0998 | 99.8 | 100.0994 |  | 1,0014 1004,30 | 1.0014 | 1001.4 | 1004.3041
| 0,1002 100,50 | 0.1002 | 100.2 | 100.5006 |  | 1,0021 1005,01 | 1.0021 | 1002.1 | 1005.0061
| 0,1001 100,40 | 0.1001 | 100.1 | 100.4003 |  | 1,0012 1004,10 | 1.0012 | 1001.2 | 1004.1035
| 0,1000 100,3 | 0.1 | 100 | 100.3 |  | 1,0015 1004,40 | 1.0015 | 1001.5 | 1004.4043
| 0,1001 100,40 | 0.1001 | 100.1 | 100.4003 |  | 1,0006 1003,50 | 1.0006 | 1000.6 | 1003.5017
| 0,1003 100,60 | 0.1003 | 100.3 | 100.6009 |  | 1,0014 1004,30 | 1.0014 | 1001.4 | 1004.3041
| 0,0998 100,10 | 0.0998 | 99.8 | 100.0994 |  | 0,9975 1000,39 | 0.9975 | 997.5 | 1000.3927
| Systematic error |  | µL | 0.4003 |  | Systematic error |  | µL | 3.953
| Systematic error |  | % | 0.3987 |  | Systematic error |  | % | 0.39375
| Random error |  | µL | 0.18312 |  | Random error |  | µL | 1.3507
| Random error |  | % | 0.18239 |  | Random error |  | % | 0.13454
| Uncertainty |  | µL | 0.76654 |  | Uncertainty |  | µL | 6.6545
4 | Andrew Alliance Picus 5-120 µL |  |  |  | 12 | Andrew Alliance Picus 50-1000 µL |  |  |  | Matrix: | Water, deionized
| Sartorius Optifit Tip 0,5-200 µL |  |  |  |  | Sartorius Optifit Tip 10-1000 µL |  |  |  |  | Temperature, Water: 20 °C
|  |  | 100 | 100.3% |  |  |  | 1000 | 99.99% |  | Air pressure 1018 hPa
| Average |  |  | 100.3201 |  | Average |  |  | 999.8913 |  | Air moisture: 31 %
| Standard deviation |  |  | 0.1233 |  | Standard deviation |  |  | 0.90447 | Z-factor | 1.0029 | mg/µL
|  | [g] | [mg] | [µL] |  |  | [g] | [mg] | [µL]
| 0,1002 100,50 | 0.1002 | 100.2 | 100.5006 |  | 0,9976 1000,49 | 0.9976 | 997.6 | 1000.493 | Technique: | Reverse
| 0,1001 100,40 | 0.1001 | 100.1 | 100.4003 |  | 0,9981 1000,99 | 0.9981 | 998.1 | 1000.9945 | Speed: | Fast
| 0,1000 100,3 | 0.1 | 100 | 100.3 |  | 0,9981 1000,99 | 0.9981 | 998.1 | 1000.9945
| 0,1002 100,50 | 0.1002 | 100.2 | 100.5006 |  | 0,9975 1000,39 | 0.9975 | 997.5 | 1000.3927
| 0,1000 100,3 | 0.1 | 100 | 100.3 |  | 0,9971 999,99 | 0.9971 | 997.1 | 999.9916
| 0,0998 100,10 | 0.0998 | 99.8 | 100.0994 |  | 0,9963 999,19 | 0.9963 | 996.3 | 999.1893
| 0,1000 100,3 | 0.1 | 100 | 100.3 |  | 0,9957 998,59 | 0.9957 | 995.7 | 998.5875
| 0,0999 100,20 | 0.0999 | 99.9 | 100.1997 |  | 0,9974 1000,29 | 0.9974 | 997.4 | 1000.2925
| 0,1000 100,3 | 0.1 | 100 | 100.3 |  | 0,9965 999,39 | 0.9965 | 996.5 | 999.3898
| 0,1000 100,3 | 0.1 | 100 | 100.3 |  | 0,9957 998,59 | 0.9957 | 995.7 | 998.5875
| Systematic error |  | µL | 0.32006 |  | Systematic error |  | µL | -0.1087
| Systematic error |  | % | 0.31904 |  | Systematic error |  | % | -0.010871
| Random error |  | µL | 0.1233 |  | Random error |  | µL | 0.90447
| Random error |  | % | 0.1229 |  | Random error |  | % | 0.090456
| Uncertainty |  | µL | 0.56665 |  | Uncertainty |  | µL | 1.9176
5 | Andrew Alliance Picus 5-120 µL |  |  |  | 13 | Andrew Alliance Picus 50-1000 µL |  |  |  | Matrix: | Water, deionized
| Sartorius Optifit Tip 0,5-200 µL |  |  |  |  | Sartorius Optifit Tip 10-1000 µL |  |  |  |  | Temperature, Water: 20 °C
| Low Retention |  | 100 | 100.4% |  | Low Retention |  | 1000 | 100.5% |  | Air pressure 1018 hPa
| Average |  |  | 100.4204 |  | Average |  |  | 1005.0161 |  | Air moisture: 31 %
| Standard deviation |  |  | 0.063435 |  | Standard deviation |  |  | 0.54204 | Z-factor | 1.0029 | mg/µL
|  | [g] | [mg] | [µL] |  |  | [g] | [mg] | [µL]
| 0,1000 100,3 | 0.1 | 100 | 100.3 |  | 1,0010 1003,90 | 1.001 | 1001 | 1003.9029 | Technique: | Forward
| 0,1001 100,40 | 0.1001 | 100.1 | 100.4003 |  | 1,0019 1004,81 | 1.0019 | 1001.9 | 1004.8055 | Speed: | Slow
| 0,1001 100,40 | 0.1001 | 100.1 | 100.4003 |  | 1,0027 1005,61 | 1.0027 | 1002.7 | 1005.6078
| 0,1002 100,50 | 0.1002 | 100.2 | 100.5006 |  | 1,0017 1004,60 | 1.0017 | 1001.7 | 1004.6049
| 0,1002 100,50 | 0.1002 | 100.2 | 100.5006 |  | 1,0017 1004,60 | 1.0017 | 1001.7 | 1004.6049
| 0,1001 100,40 | 0.1001 | 100.1 | 100.4003 |  | 1,0023 1005,21 | 1.0023 | 1002.3 | 1005.2067
| 0,1001 100,40 | 0.1001 | 100.1 | 100.4003 |  | 1,0026 1005,51 | 1.0026 | 1002.6 | 1005.5075
| 0,1002 100,50 | 0.1002 | 100.2 | 100.5006 |  | 1,0026 1005,51 | 1.0026 | 1002.6 | 1005.5075
| 0,1001 100,40 | 0.1001 | 100.1 | 100.4003 |  | 1,0021 1005,01 | 1.0021 | 1002.1 | 1005.0061
| 0,1001 100,40 | 0.1001 | 100.1 | 100.4003 |  | 1,0025 1005,41 | 1.0025 | 1002.5 | 1005.4072
| Systematic error |  | µL | 0.42036 |  | Systematic error |  | µL | 5.0161
| Systematic error |  | % | 0.4186 |  | Systematic error |  | % | 0.49911
| Random error |  | µL | 0.063435 |  | Random error |  | µL | 0.54204
| Random error |  | % | 0.06317 |  | Random error |  | % | 0.053934
| Uncertainty |  | µL | 0.54723 |  | Uncertainty |  | µL | 6.1002
6 | Andrew Alliance Picus 5-120 µL |  |  |  | 14 | Andrew Alliance Picus 50-1000 µL |  |  |  | Matrix: | Water, deionized
| Sartorius Optifit Tip 0,5-200 µL |  |  |  |  | Sartorius Optifit Tip 10-1000 µL |  |  |  |  | Temperature, Water: 20 °C
| Low Retention |  | 100 | 100% |  | Low Retention |  | 1000 | 100.5% |  | Air pressure 1017 hPa
| Average |  |  | 100.0492 |  | Average |  |  | 1004.7855 |  | Air moisture: 43 %
| Standard deviation |  |  | 0.085239 |  | Standard deviation |  |  | 0.4723 | Z-factor | 1.0029 | mg/µL
|  | [g] | [mg] | [µL] |  |  | [g] | [mg] | [µL]
| 0,0996 99,90 | 0.0996 | 99.6 | 99.8988 |  | 1,0008 1003,70 | 1.0008 | 1000.8 | 1003.7023 | Technique: | Forward
| 0,0997 100,00 | 0.0997 | 99.7 | 99.9991 |  | 1,0024 1005,31 | 1.0024 | 1002.4 | 1005.307 | Speed: | Fast
| 0,0999 100,20 | 0.0999 | 99.9 | 100.1997 |  | 1,0018 1004,71 | 1.0018 | 1001.8 | 1004.7052
| 0,0997 100,00 | 0.0997 | 99.7 | 99.9991 |  | 1,0023 1005,21 | 1.0023 | 1002.3 | 1005.2067
| 0,0997 100,00 | 0.0997 | 99.7 | 99.9991 |  | 1,0016 1004,50 | 1.0016 | 1001.6 | 1004.5046
| 0,0997 100,00 | 0.0997 | 99.7 | 99.9991 |  | 1,0022 1005,11 | 1.0022 | 1002.2 | 1005.1064
| 0,0998 100,10 | 0.0998 | 99.8 | 100.0994 |  | 1,0020 1004,91 | 1.002 | 1002 | 1004.9058
| 0,0998 100,10 | 0.0998 | 99.8 | 100.0994 |  | 1,0022 1005,11 | 1.0022 | 1002.2 | 1005.1064
| 0,0998 100,10 | 0.0998 | 99.8 | 100.0994 |  | 1,0019 1004,81 | 1.0019 | 1001.9 | 1004.8055
| 0,0998 100,10 | 0.0998 | 99.8 | 100.0994 |  | 1,0016 1004,50 | 1.0016 | 1001.6 | 1004.5046
| Systematic error |  | µL | 0.04925 |  | Systematic error |  | µL | 4.7855
| Systematic error |  | % | 0.049226 |  | Systematic error |  | % | 0.47627
| Random error |  | µL | 0.085239 |  | Random error |  | µL | 0.4723
| Random error |  | % | 0.085197 |  | Random error |  | % | 0.047005
| Uncertainty |  | µL | 0.21973 |  | Uncertainty |  | µL | 5.73
7 | Andrew Alliance Picus 5-120 µL |  |  |  | 15 | Andrew Alliance Picus 50-1000 µL |  |  |  | Matrix: | Water, deionized
| Sartorius Optifit Tip 0,5-200 µL |  |  |  |  | Sartorius Optifit Tip 10-1000 µL |  |  |  |  | Temperature, Water: 20 °C
| Low Retention |  | 100 | 100.2% |  | Low Retention |  | 1000 | 100.4% |  | Air pressure 1017 hPa
| Average |  |  | 100.2498 |  | Average |  |  | 1004.1436 |  | Air moisture: 43 %
| Standard deviation |  |  | 0.10834 |  | Standard deviation |  |  | 0.93029 | Z-factor | 1.0029 | mg/µL
|  | [g] | [mg] | [µL] |  |  | [g] | [mg] | [µL]
| 0,0999 100,20 | 0.0999 | 99.9 | 100.1997 |  | 1,0031 1006,01 | 1.0031 | 1003.1 | 1006.009 | Technique: | Reverse
| 0,1001 100,40 | 0.1001 | 100.1 | 100.4003 |  | 1,0017 1004,60 | 1.0017 | 1001.7 | 1004.6049 | Speed: | Slow
| 0,0997 100,00 | 0.0997 | 99.7 | 99.9991 |  | 1,0011 1004,00 | 1.0011 | 1001.1 | 1004.0032
| 0,0999 100,20 | 0.0999 | 99.9 | 100.1997 |  | 1,0016 1004,50 | 1.0016 | 1001.6 | 1004.5046
| 0,1000 100,3 | 0.1 | 100 | 100.3 |  | 1,0013 1004,20 | 1.0013 | 1001.3 | 1004.2038
| 0,1000 100,3 | 0.1 | 100 | 100.3 |  | 1,0019 1004,81 | 1.0019 | 1001.9 | 1004.8055
| 0,1000 100,3 | 0.1 | 100 | 100.3 |  | 1,0004 1003,30 | 1.0004 | 1000.4 | 1003.3012
| 0,1000 100,3 | 0.1 | 100 | 100.3 |  | 1,0009 1003,80 | 1.0009 | 1000.9 | 1003.8026
| 0,1000 100,3 | 0.1 | 100 | 100.3 |  | 1,0007 1003,60 | 1.0007 | 1000.7 | 1003.602
| 0,0999 100,20 | 0.0999 | 99.9 | 100.1997 |  | 0,9997 1002,60 | 0.9997 | 999.7 | 1002.5991
| Systematic error |  | µL | 0.24985 |  | Systematic error |  | µL | 4.1436
| Systematic error |  | % | 0.24923 |  | Systematic error |  | % | 0.41265
| Random error |  | µL | 0.10834 |  | Random error |  | µL | 0.93029
| Random error |  | % | 0.10807 |  | Random error |  | % | 0.092645
| Uncertainty |  | µL | 0.46652 |  | Uncertainty |  | µL | 6.0042
8 | Andrew Alliance Picus 5-120 µL |  |  |  | 16 | Andrew Alliance Picus 50-1000 µL |  |  |  | Matrix: | Water, deionized
| Sartorius Optifit Tip 0,5-200 µL |  |  |  |  | Sartorius Optifit Tip 10-1000 µL |  |  |  |  | Temperature, Water: 20 °C
| Low Retention |  | 100 | 100.4% |  | Low Retention |  | 1000 | 100.1% |  | Air pressure 1017 hPa
| Average |  |  | 100.3602 |  | Average |  |  | 1000.8741 |  | Air moisture: 43 %
| Standard deviation |  |  | 0.10782 |  | Standard deviation |  |  | 0.76203 | Z-factor | 1.0029 | mg/µL
|  | [g] | [mg] | [µL] |  |  | [g] | [mg] | [µL]
| 0,1002 100,50 | 0.1002 | 100.2 | 100.5006 |  | 0,9988 1001,70 | 0.9988 | 998.8 | 1001.6965 | Technique: | Reverse
| 0,1001 100,40 | 0.1001 | 100.1 | 100.4003 |  | 0,9991 1002,00 | 0.9991 | 999.1 | 1001.9974 | Speed: | Fast
| 0,1001 100,40 | 0.1001 | 100.1 | 100.4003 |  | 0,9977 1000,59 | 0.9977 | 997.7 | 1000.5933
| 0,1002 100,50 | 0.1002 | 100.2 | 100.5006 |  | 0,9983 1001,20 | 0.9983 | 998.3 | 1001.1951
| 0,1001 100,40 | 0.1001 | 100.1 | 100.4003 |  | 0,9972 1000,09 | 0.9972 | 997.2 | 1000.0919
| 0,1000 100,3 | 0.1 | 100 | 100.3 |  | 0,9985 1001,40 | 0.9985 | 998.5 | 1001.3956
| 0,0999 100,20 | 0.0999 | 99.9 | 100.1997 |  | 0,9984 1001,30 | 0.9984 | 998.4 | 1001.2954
| 0,1001 100,40 | 0.1001 | 100.1 | 100.4003 |  | 0,9972 1000,09 | 0.9972 | 997.2 | 1000.0919
| 0,0999 100,20 | 0.0999 | 99.9 | 100.1997 |  | 0,9978 1000,69 | 0.9978 | 997.8 | 1000.6936
| 0,1000 100,3 | 0.1 | 100 | 100.3 |  | 0,9968 999,69 | 0.9968 | 996.8 | 999.6907
| Systematic error |  | µL | 0.36018 |  | Systematic error |  | µL | 0.87414
| Systematic error |  | % | 0.35889 |  | Systematic error |  | % | 0.087338
| Random error |  | µL | 0.10782 |  | Random error |  | µL | 0.76203
| Random error |  | % | 0.10743 |  | Random error |  | % | 0.076136
| Uncertainty |  | µL | 0.57582 |  | Uncertainty |  | µL | 2.3982
[5 empty] | 17 | Andrew Alliance Picus 50-1000 µL |  |  |  | Matrix: | Water, deionized
[6 empty] | Sartorius Optifit Tip 10-1000 µL |  |  |  |  | Temperature, Water: 20 °C
[6 empty] | Wide Bore |  | 1000 | 100.8% |  | Air pressure 1017 hPa
[6 empty] | Average |  |  | 1007.7139 |  | Air moisture: 43 %
[6 empty] | Standard deviation |  |  | 0.45837 | Z-factor | 1.0029 | mg/µL
[7 empty] | [g] | [mg] | [µL]
[6 empty] | 1,0045 1007,41 | 1.0045 | 1004.5 | 1007.413 | Technique: | Forward
[6 empty] | 1,0048 1007,71 | 1.0048 | 1004.8 | 1007.7139 | Speed: | Slow
[6 empty] | 1,0052 1008,12 | 1.0052 | 1005.2 | 1008.1151
[6 empty] | 1,0042 1007,11 | 1.0042 | 1004.2 | 1007.1122
[6 empty] | 1,0046 1007,51 | 1.0046 | 1004.6 | 1007.5133
[6 empty] | 1,0057 1008,62 | 1.0057 | 1005.7 | 1008.6165
[6 empty] | 1,0048 1007,71 | 1.0048 | 1004.8 | 1007.7139
[6 empty] | 1,0052 1008,12 | 1.0052 | 1005.2 | 1008.1151
[6 empty] | 1,0047 1007,61 | 1.0047 | 1004.7 | 1007.6136
[6 empty] | 1,0043 1007,21 | 1.0043 | 1004.3 | 1007.2125
[6 empty] | Systematic error |  | µL | 7.7139
[6 empty] | Systematic error |  | % | 0.76549
[6 empty] | Random error |  | µL | 0.45837
[6 empty] | Random error |  | % | 0.045486
[6 empty] | Uncertainty |  | µL | 8.6307
[5 empty] | 18 | Andrew Alliance Picus 50-1000 µL |  |  |  | Matrix: | Water, deionized
[6 empty] | Sartorius Optifit Tip 10-1000 µL |  |  |  |  | Temperature, Water: 20 °C
[6 empty] | Wide Bore |  | 1000 | 100.4% |  | Air pressure 1017 hPa
[6 empty] | Average |  |  | 1004.1235 |  | Air moisture: 43 %
[6 empty] | Standard deviation |  |  | 0.75466 | Z-factor | 1.0029 | mg/µL
[7 empty] | [g] | [mg] | [µL]
[6 empty] | 0,9996 1002,50 | 0.9996 | 999.6 | 1002.4988 | Technique: | Forward
[6 empty] | 1,0021 1005,01 | 1.0021 | 1002.1 | 1005.0061 | Speed: | Fast
[6 empty] | 1,0018 1004,71 | 1.0018 | 1001.8 | 1004.7052
[6 empty] | 1,0017 1004,60 | 1.0017 | 1001.7 | 1004.6049
[6 empty] | 1,0019 1004,81 | 1.0019 | 1001.9 | 1004.8055
[6 empty] | 1,0011 1004,00 | 1.0011 | 1001.1 | 1004.0032
[6 empty] | 1,0015 1004,40 | 1.0015 | 1001.5 | 1004.4043
[6 empty] | 1,0006 1003,50 | 1.0006 | 1000.6 | 1003.5017
[6 empty] | 1,0011 1004,00 | 1.0011 | 1001.1 | 1004.0032
[6 empty] | 1,0008 1003,70 | 1.0008 | 1000.8 | 1003.7023
[6 empty] | Systematic error |  | µL | 4.1235
[6 empty] | Systematic error |  | % | 0.41066
[6 empty] | Random error |  | µL | 0.75466
[6 empty] | Random error |  | % | 0.075156
[6 empty] | Uncertainty |  | µL | 5.6329
[5 empty] | 19 | Andrew Alliance Picus 50-1000 µL |  |  |  | Matrix: | Water, deionized
[6 empty] | Sartorius Optifit Tip 10-1000 µL |  |  |  |  | Temperature, Water: 20 °C
[6 empty] | Wide Bore |  | 1000 | 100.4% |  | Air pressure 1017 hPa
[6 empty] | Average |  |  | 1004.3442 |  | Air moisture: 43 %
[6 empty] | Standard deviation |  |  | 1.0371 | Z-factor | 1.0029 | mg/µL
[7 empty] | [g] | [mg] | [µL]
[6 empty] | 1,0037 1006,61 | 1.0037 | 1003.7 | 1006.6107 | Technique: | Reverse
[6 empty] | 1,0026 1005,51 | 1.0026 | 1002.6 | 1005.5075 | Speed: | Slow
[6 empty] | 1,0000 1002,9 | 1 | 1000 | 1002.9
[6 empty] | 1,0012 1004,10 | 1.0012 | 1001.2 | 1004.1035
[6 empty] | 1,0010 1003,90 | 1.001 | 1001 | 1003.9029
[6 empty] | 1,0011 1004,00 | 1.0011 | 1001.1 | 1004.0032
[6 empty] | 1,0006 1003,50 | 1.0006 | 1000.6 | 1003.5017
[6 empty] | 1,0015 1004,40 | 1.0015 | 1001.5 | 1004.4043
[6 empty] | 1,0014 1004,30 | 1.0014 | 1001.4 | 1004.3041
[6 empty] | 1,0013 1004,20 | 1.0013 | 1001.3 | 1004.2038
[6 empty] | Systematic error |  | µL | 4.3442
[6 empty] | Systematic error |  | % | 0.43254
[6 empty] | Random error |  | µL | 1.0371
[6 empty] | Random error |  | % | 0.10326
[6 empty] | Uncertainty |  | µL | 6.4183
[5 empty] | 20 | Andrew Alliance Picus 50-1000 µL |  |  |  | Matrix: | Water, deionized
[6 empty] | Sartorius Optifit Tip 10-1000 µL |  |  |  |  | Temperature, Water: 20 °C
[6 empty] | Wide Bore |  | 1000 | 100% |  | Air pressure 1017 hPa
[6 empty] | Average |  |  | 1000.0417 |  | Air moisture: 43 %
[6 empty] | Standard deviation |  |  | 0.45654 | Z-factor | 1.0029 | mg/µL
[7 empty] | [g] | [mg] | [µL]
[6 empty] | 0,9979 1000,79 | 0.9979 | 997.9 | 1000.7939 | Technique: | Reverse
[6 empty] | 0,9976 1000,49 | 0.9976 | 997.6 | 1000.493 | Speed: | Fast
[6 empty] | 0,9975 1000,39 | 0.9975 | 997.5 | 1000.3927
[6 empty] | 0,9970 999,89 | 0.997 | 997 | 999.8913
[6 empty] | 0,9970 999,89 | 0.997 | 997 | 999.8913
[6 empty] | 0,9972 1000,09 | 0.9972 | 997.2 | 1000.0919
[6 empty] | 0,9972 1000,09 | 0.9972 | 997.2 | 1000.0919
[6 empty] | 0,9971 999,99 | 0.9971 | 997.1 | 999.9916
[6 empty] | 0,9967 999,59 | 0.9967 | 996.7 | 999.5904
[6 empty] | 0,9963 999,19 | 0.9963 | 996.3 | 999.1893
[6 empty] | Systematic error |  | µL | 0.041735
[6 empty] | Systematic error |  | % | 0.0041733
[6 empty] | Random error |  | µL | 0.45654
[6 empty] | Random error |  | % | 0.045652
[6 empty] | Uncertainty |  | µL | 0.95481

## Blood
1 | Andrew Alliance Picus 5-120 µL |  |  |  | 9 | Andrew Alliance Picus 50-1000 µL |  |  |  | Matrix: | Blood (Sheep, +citrate)
| Sartorius Optifit Tip 0,5-200 µL |  |  |  |  | Sartorius Optifit Tip 10-1000 µL |  |  |  |  | Bio Trading BTSN250
|  |  | 100 | 100.5% |  |  |  | 1000 | 101.3% |  | Lot: #16130 SN78
| Average |  |  | 100.4506 |  | Average |  |  | 1012.5216 |  | MW:
| Standard deviation |  |  | 1.1452 |  | Standard deviation |  |  | 3.2542 | Density (experimental): | 1.043 | mg/µL
|  | [g] | [mg] | [µL] |  |  | [g] | [mg] | [µL] | Density (literature): | 1.0275 | mg/µL
| 0,1037 104,00 | 0.1037 | 103.7 | 99.4247 |  | 1,0512 1054,25 | 1.0512 | 1051.2 | 1007.8619
| 0,1066 106,91 | 0.1066 | 106.6 | 102.2052 |  | 1,0596 1062,67 | 1.0596 | 1059.6 | 1015.9156 | Technique: | Forward
| 0,1044 104,70 | 0.1044 | 104.4 | 100.0959 |  | 1,0562 1059,26 | 1.0562 | 1056.2 | 1012.6558 | Speed: | Slow
| 0,1045 104,80 | 0.1045 | 104.5 | 100.1918 |  | 1,0583 1061,37 | 1.0583 | 1058.3 | 1014.6692
| 0,1042 104,50 | 0.1042 | 104.2 | 99.9041 |  | 1,0510 1054,05 | 1.051 | 1051 | 1007.6702
| 0,1029 103,20 | 0.1029 | 102.9 | 98.6577 |  | 1,0539 1056,96 | 1.0539 | 1053.9 | 1010.4506
| 0,1052 105,51 | 0.1052 | 105.2 | 100.8629 |  | 1,0558 1058,86 | 1.0558 | 1055.8 | 1012.2723
| 0,1041 104,40 | 0.1041 | 104.1 | 99.8082 |  | 1,0547 1057,76 | 1.0547 | 1054.7 | 1011.2176
| 0,1056 105,91 | 0.1056 | 105.6 | 101.2464 |  | 1,0598 1062,87 | 1.0598 | 1059.8 | 1016.1074
| 0,1065 106,81 | 0.1065 | 106.5 | 102.1093 |  | 1,0601 1063,17 | 1.0601 | 1060.1 | 1016.395
| Systematic error |  | µL | 0.45062 |  | Systematic error |  | µL | 12.5216
| Systematic error |  | % | 0.4486 |  | Systematic error |  | % | 1.2367
| Random error |  | µL | 1.1452 |  | Random error |  | µL | 3.2542
| Random error |  | % | 1.1401 |  | Random error |  | % | 0.3214
| Uncertainty |  | µL | 2.7411 |  | Uncertainty |  | µL | 19.0301
2 | Andrew Alliance Picus 5-120 µL |  |  |  | 10 | Andrew Alliance Picus 50-1000 µL |  |  |  | Matrix: | Blood (Sheep, +sitraatti)
| Sartorius Optifit Tip 0,5-200 µL |  |  |  |  | Sartorius Optifit Tip 10-1000 µL |  |  |  |  | Bio Trading BTSN250
|  |  | 100 | 98.09% |  |  |  | 1000 | 101.1% |  | Lot: #16130 SN78
| Average |  |  | 98.092 |  | Average |  |  | 1010.767 |  | MW:
| Standard deviation |  |  | 0.55805 |  | Standard deviation |  |  | 1.2887 | Density (experimental): | 1.043 | mg/µL
|  | [g] | [mg] | [µL] |  |  | [g] | [mg] | [µL] | Density (literature): | 1.0275 | mg/µL
| 0,1029 103,20 | 0.1029 | 102.9 | 98.6577 |  | 1,0540 1057,06 | 1.054 | 1054 | 1010.5465
| 0,1016 101,89 | 0.1016 | 101.6 | 97.4113 |  | 1,0547 1057,76 | 1.0547 | 1054.7 | 1011.2176 | Technique: | Forward
| 0,1021 102,40 | 0.1021 | 102.1 | 97.8907 |  | 1,0532 1056,25 | 1.0532 | 1053.2 | 1009.7795 | Speed: | Fast
| 0,1018 102,10 | 0.1018 | 101.8 | 97.6031 |  | 1,0562 1059,26 | 1.0562 | 1056.2 | 1012.6558
| 0,1027 103,00 | 0.1027 | 102.7 | 98.466 |  | 1,0544 1057,46 | 1.0544 | 1054.4 | 1010.93
| 0,1018 102,10 | 0.1018 | 101.8 | 97.6031 |  | 1,0549 1057,96 | 1.0549 | 1054.9 | 1011.4094
| 0,1017 101,99 | 0.1017 | 101.7 | 97.5072 |  | 1,0514 1054,45 | 1.0514 | 1051.4 | 1008.0537
| 0,1031 103,40 | 0.1031 | 103.1 | 98.8495 |  | 1,0548 1057,86 | 1.0548 | 1054.8 | 1011.3135
| 0,1030 103,30 | 0.103 | 103 | 98.7536 |  | 1,0533 1056,35 | 1.0533 | 1053.3 | 1009.8754
| 0,1024 102,70 | 0.1024 | 102.4 | 98.1783 |  | 1,0554 1058,46 | 1.0554 | 1055.4 | 1011.8888
| Systematic error |  | µL | -1.908 |  | Systematic error |  | µL | 10.767
| Systematic error |  | % | -1.9451 |  | Systematic error |  | % | 1.0652
| Random error |  | µL | 0.55805 |  | Random error |  | µL | 1.2887
| Random error |  | % | 0.5689 |  | Random error |  | % | 0.1275
| Uncertainty |  | µL | 3.0241 |  | Uncertainty |  | µL | 13.3445
3 | Andrew Alliance Picus 5-120 µL |  |  | 2016-05-31 | 11 | Andrew Alliance Picus 50-1000 µL |  |  | 2016-05-31 | Matrix: | Blood (Sheep, +sitraatti)
| Sartorius Optifit Tip 0,5-200 µL |  |  |  |  | Sartorius Optifit Tip 10-1000 µL |  |  |  |  | Bio Trading BTSN250
|  |  | 100 | 99.89% |  |  |  | 1000 | 100.4% |  | Lot: #16130 SN78
| Average |  |  | 99.8945 |  | Average |  |  | 1003.72 |  | MW:
| Standard deviation |  |  | 0.64308 |  | Standard deviation |  |  | 2.9447 | Density (experimental): | 1.043 | mg/µL
|  | [g] | [mg] | [µL] |  |  | [g] | [mg] | [µL] | Density (literature): | 1.0275 | mg/µL
| 0,1046 104,90 | 0.1046 | 104.6 | 100.2876 |  | 1,0491 1052,14 | 1.0491 | 1049.1 | 1005.8485
| 0,1046 104,90 | 0.1046 | 104.6 | 100.2876 |  | 1,0438 1046,83 | 1.0438 | 1043.8 | 1000.767 | Technique: | Reverse
| 0,1043 104,60 | 0.1043 | 104.3 | 100 |  | 1,0418 1044,82 | 1.0418 | 1041.8 | 998.8495 | Speed: | Slow
| 0,1049 105,20 | 0.1049 | 104.9 | 100.5753 |  | 1,0437 1046,73 | 1.0437 | 1043.7 | 1000.6711
| 0,1042 104,50 | 0.1042 | 104.2 | 99.9041 |  | 1,0475 1050,54 | 1.0475 | 1047.5 | 1004.3145
| 0,1045 104,80 | 0.1045 | 104.5 | 100.1918 |  | 1,0450 1048,03 | 1.045 | 1045 | 1001.9175
| 0,1044 104,70 | 0.1044 | 104.4 | 100.0959 |  | 1,0496 1052,64 | 1.0496 | 1049.6 | 1006.3279
| 0,1044 104,70 | 0.1044 | 104.4 | 100.0959 |  | 1,0480 1051,04 | 1.048 | 1048 | 1004.7939
| 0,1033 103,60 | 0.1033 | 103.3 | 99.0412 |  | 1,0496 1052,64 | 1.0496 | 1049.6 | 1006.3279
| 0,1027 103,00 | 0.1027 | 102.7 | 98.466 |  | 1,0507 1053,75 | 1.0507 | 1050.7 | 1007.3826
| Systematic error |  | µL | -0.10547 |  | Systematic error |  | µL | 3.72
| Systematic error |  | % | -0.10558 |  | Systematic error |  | % | 0.37063
| Random error |  | µL | 0.64308 |  | Random error |  | µL | 2.9447
| Random error |  | % | 0.64376 |  | Random error |  | % | 0.29338
| Uncertainty |  | µL | 1.3916 |  | Uncertainty |  | µL | 9.6094
4 | Andrew Alliance Picus 5-120 µL |  |  | 2016-05-31 | 12 | Andrew Alliance Picus 50-1000 µL |  |  | 2016-05-31 | Matrix: | Blood (Sheep, +sitraatti)
| Sartorius Optifit Tip 0,5-200 µL |  |  |  |  | Sartorius Optifit Tip 10-1000 µL |  |  |  |  | Bio Trading BTSN250
|  |  | 100 | 99.85% |  |  |  | 1000 | 99.61% |  | Lot: #16130 SN78
| Average |  |  | 99.8466 |  | Average |  |  | 996.117 |  | MW:
| Standard deviation |  |  | 0.6363 |  | Standard deviation |  |  | 4.6641 | Density (experimental): | 1.043 | mg/µL
|  | [g] | [mg] | [µL] |  |  | [g] | [mg] | [µL] | Density (literature): | 1.0275 | mg/µL
| 0,1041 104,40 | 0.1041 | 104.1 | 99.8082 |  | 1,0360 1039,00 | 1.036 | 1036 | 993.2886
| 0,1039 104,20 | 0.1039 | 103.9 | 99.6165 |  | 1,0266 1029,58 | 1.0266 | 1026.6 | 984.2761 | Technique: | Reverse
| 0,1043 104,60 | 0.1043 | 104.3 | 100 |  | 1,0407 1043,72 | 1.0407 | 1040.7 | 997.7948 | Speed: | Fast
| 0,1044 104,70 | 0.1044 | 104.4 | 100.0959 |  | 1,0427 1045,72 | 1.0427 | 1042.7 | 999.7124
| 0,1027 103,00 | 0.1027 | 102.7 | 98.466 |  | 1,0406 1043,62 | 1.0406 | 1040.6 | 997.6989
| 0,1048 105,10 | 0.1048 | 104.8 | 100.4794 |  | 1,0374 1040,41 | 1.0374 | 1037.4 | 994.6309
| 0,1041 104,40 | 0.1041 | 104.1 | 99.8082 |  | 1,0399 1042,92 | 1.0399 | 1039.9 | 997.0278
| 0,1049 105,20 | 0.1049 | 104.9 | 100.5753 |  | 1,0411 1044,12 | 1.0411 | 1041.1 | 998.1783
| 0,1035 103,80 | 0.1035 | 103.5 | 99.233 |  | 1,0412 1044,22 | 1.0412 | 1041.2 | 998.2742
| 0,1047 105,00 | 0.1047 | 104.7 | 100.3835 |  | 1,0433 1046,33 | 1.0433 | 1043.3 | 1000.2876
| Systematic error |  | µL | -0.1534 |  | Systematic error |  | µL | -3.883
| Systematic error |  | % | -0.15364 |  | Systematic error |  | % | -0.38982
| Random error |  | µL | 0.6363 |  | Random error |  | µL | 4.6641
| Random error |  | % | 0.63728 |  | Random error |  | % | 0.46823
| Uncertainty |  | µL | 1.426 |  | Uncertainty |  | µL | 13.2113
5 | Andrew Alliance Picus 5-120 µL |  |  | 2016-05-31 | 13 | Andrew Alliance Picus 50-1000 µL |  |  | 2016-05-31 | Matrix: | Blood (Sheep, +sitraatti)
| Sartorius Optifit Tip 0,5-200 µL |  |  |  |  | Sartorius Optifit Tip 10-1000 µL |  |  |  |  | Bio Trading BTSN250
| Low Retention |  | 100 | 102.2% |  | Low Retention |  | 1000 | 100.3% |  | Lot: #16130 SN78
| Average |  |  | 102.2052 |  | Average |  |  | 1003.0105 |  | MW:
| Standard deviation |  |  | 1.5314 |  | Standard deviation |  |  | 2.3429 | Density (experimental): | 1.043 | mg/µL
|  | [g] | [mg] | [µL] |  |  | [g] | [mg] | [µL] | Density (literature): | 1.0275 | mg/µL | < Sheep (vain plasma?)
| 0,1039 104,20 | 0.1039 | 103.9 | 99.6165 |  | 1,0432 1046,23 | 1.0432 | 1043.2 | 1000.1918
| 0,1058 106,11 | 0.1058 | 105.8 | 101.4382 |  | 1,0481 1051,14 | 1.0481 | 1048.1 | 1004.8897 | Technique: | Forward
| 0,1060 106,31 | 0.106 | 106 | 101.6299 |  | 1,0469 1049,94 | 1.0469 | 1046.9 | 1003.7392 | Speed: | Slow
| 0,1066 106,91 | 0.1066 | 106.6 | 102.2052 |  | 1,0477 1050,74 | 1.0477 | 1047.7 | 1004.5062
| 0,1072 107,51 | 0.1072 | 107.2 | 102.7804 |  | 1,0446 1047,63 | 1.0446 | 1044.6 | 1001.534
| 0,1073 107,61 | 0.1073 | 107.3 | 102.8763 |  | 1,0469 1049,94 | 1.0469 | 1046.9 | 1003.7392
| 0,1068 107,11 | 0.1068 | 106.8 | 102.3969 |  | 1,0476 1050,64 | 1.0476 | 1047.6 | 1004.4104
| 0,1047 105,00 | 0.1047 | 104.7 | 100.3835 |  | 1,0412 1044,22 | 1.0412 | 1041.2 | 998.2742
| 0,1087 109,02 | 0.1087 | 108.7 | 104.2186 |  | 1,0461 1049,13 | 1.0461 | 1046.1 | 1002.9722
| 0,1090 109,32 | 0.109 | 109 | 104.5062 |  | 1,0491 1052,14 | 1.0491 | 1049.1 | 1005.8485
| Systematic error |  | µL | 2.2052 |  | Systematic error |  | µL | 3.0105
| Systematic error |  | % | 2.1576 |  | Systematic error |  | % | 0.30015
| Random error |  | µL | 1.5314 |  | Random error |  | µL | 2.3429
| Random error |  | % | 1.4983 |  | Random error |  | % | 0.23359
| Uncertainty |  | µL | 5.2679 |  | Uncertainty |  | µL | 7.6964
6 | Andrew Alliance Picus 5-120 µL |  |  | 2016-05-31 | 14 | Andrew Alliance Picus 50-1000 µL |  |  | 2016-05-31 | Matrix: | Blood (Sheep, +sitraatti)
| Sartorius Optifit Tip 0,5-200 µL |  |  |  |  | Sartorius Optifit Tip 10-1000 µL |  |  |  |  | Bio Trading BTSN250
| Low Retention |  | 100 | 99.63% |  | Low Retention |  | 1000 | 100.8% |  | Lot: #16130 SN78
| Average |  |  | 99.6261 |  | Average |  |  | 1008.092 |  | MW:
| Standard deviation |  |  | 0.98758 |  | Standard deviation |  |  | 1.4172 | Density (experimental): | 1.043 | mg/µL
|  | [g] | [mg] | [µL] |  |  | [g] | [mg] | [µL] | Density (literature): | 1.0275 | mg/µL
| 0,1022 102,50 | 0.1022 | 102.2 | 97.9866 |  | 1,0504 1053,45 | 1.0504 | 1050.4 | 1007.0949
| 0,1033 103,60 | 0.1033 | 103.3 | 99.0412 |  | 1,0516 1054,65 | 1.0516 | 1051.6 | 1008.2454 | Technique: | Forward
| 0,1027 103,00 | 0.1027 | 102.7 | 98.466 |  | 1,0486 1051,64 | 1.0486 | 1048.6 | 1005.3691 | Speed: | Fast
| 0,1039 104,20 | 0.1039 | 103.9 | 99.6165 |  | 1,0507 1053,75 | 1.0507 | 1050.7 | 1007.3826
| 0,1037 104,00 | 0.1037 | 103.7 | 99.4247 |  | 1,0534 1056,45 | 1.0534 | 1053.4 | 1009.9712
| 0,1038 104,10 | 0.1038 | 103.8 | 99.5206 |  | 1,0507 1053,75 | 1.0507 | 1050.7 | 1007.3826
| 0,1041 104,40 | 0.1041 | 104.1 | 99.8082 |  | 1,0520 1055,05 | 1.052 | 1052 | 1008.629
| 0,1050 105,30 | 0.105 | 105 | 100.6711 |  | 1,0533 1056,35 | 1.0533 | 1053.3 | 1009.8754
| 0,1049 105,20 | 0.1049 | 104.9 | 100.5753 |  | 1,0527 1055,75 | 1.0527 | 1052.7 | 1009.3001
| 0,1055 105,81 | 0.1055 | 105.5 | 101.1505 |  | 1,0510 1054,05 | 1.051 | 1051 | 1007.6702
| Systematic error |  | µL | -0.37392 |  | Systematic error |  | µL | 8.092
| Systematic error |  | % | -0.37532 |  | Systematic error |  | % | 0.80271
| Random error |  | µL | 0.98758 |  | Random error |  | µL | 1.4172
| Random error |  | % | 0.99129 |  | Random error |  | % | 0.14058
| Uncertainty |  | µL | 2.3491 |  | Uncertainty |  | µL | 10.9264
7 | Andrew Alliance Picus 5-120 µL |  |  | 2016-05-31 | 15 | Andrew Alliance Picus 50-1000 µL |  |  | 2016-05-31 | Matrix: | Blood (Sheep, +sitraatti)
| Sartorius Optifit Tip 0,5-200 µL |  |  |  |  | Sartorius Optifit Tip 10-1000 µL |  |  |  |  | Bio Trading BTSN250
| Low Retention |  | 100 | 101.7% |  | Low Retention |  | 1000 | 99.27% |  | Lot: #16130 SN78
| Average |  |  | 101.7354 |  | Average |  |  | 992.7421 |  | MW:
| Standard deviation |  |  | 1.5317 |  | Standard deviation |  |  | 1.4583 | Density (experimental): | 1.043 | mg/µL
|  | [g] | [mg] | [µL] |  |  | [g] | [mg] | [µL] | Density (literature): | 1.0275 | mg/µL
| 0,1047 105,00 | 0.1047 | 104.7 | 100.3835 |  | 1,0342 1037,20 | 1.0342 | 1034.2 | 991.5628
| 0,1042 104,50 | 0.1042 | 104.2 | 99.9041 |  | 1,0331 1036,10 | 1.0331 | 1033.1 | 990.5081 | Technique: | Reverse
| 0,1050 105,30 | 0.105 | 105 | 100.6711 |  | 1,0337 1036,70 | 1.0337 | 1033.7 | 991.0834 | Speed: | Slow
| 0,1079 108,21 | 0.1079 | 107.9 | 103.4516 |  | 1,0355 1038,50 | 1.0355 | 1035.5 | 992.8092
| 0,1064 106,71 | 0.1064 | 106.4 | 102.0134 |  | 1,0353 1038,30 | 1.0353 | 1035.3 | 992.6174
| 0,1065 106,81 | 0.1065 | 106.5 | 102.1093 |  | 1,0365 1039,51 | 1.0365 | 1036.5 | 993.768
| 0,1069 107,21 | 0.1069 | 106.9 | 102.4928 |  | 1,0372 1040,21 | 1.0372 | 1037.2 | 994.4391
| 0,1063 106,61 | 0.1063 | 106.3 | 101.9175 |  | 1,0346 1037,60 | 1.0346 | 1034.6 | 991.9463
| 0,1090 109,32 | 0.109 | 109 | 104.5062 |  | 1,0367 1039,71 | 1.0367 | 1036.7 | 993.9597
| 0,1042 104,50 | 0.1042 | 104.2 | 99.9041 |  | 1,0375 1040,51 | 1.0375 | 1037.5 | 994.7267
| Systematic error |  | µL | 1.7354 |  | Systematic error |  | µL | -7.2579
| Systematic error |  | % | 1.7058 |  | Systematic error |  | % | -0.7311
| Random error |  | µL | 1.5317 |  | Random error |  | µL | 1.4583
| Random error |  | % | 1.5055 |  | Random error |  | % | 0.1469
| Uncertainty |  | µL | 4.7987 |  | Uncertainty |  | µL | 10.1745
8 | Andrew Alliance Picus 5-120 µL |  |  | 2016-05-31 | 16 | Andrew Alliance Picus 50-1000 µL |  |  | 2016-05-31 | Matrix: | Blood (Sheep, +sitraatti)
| Sartorius Optifit Tip 0,5-200 µL |  |  |  |  | Sartorius Optifit Tip 10-1000 µL |  |  |  |  | Bio Trading BTSN250
| Low Retention |  | 100 | 99.06% |  | Low Retention |  | 1000 | 99.24% |  | Lot: #16130 SN78
| Average |  |  | 99.0604 |  | Average |  |  | 992.4161 |  | MW:
| Standard deviation |  |  | 1.3176 |  | Standard deviation |  |  | 1.9395 | Density (experimental): | 1.043 | mg/µL
|  | [g] | [mg] | [µL] |  |  | [g] | [mg] | [µL] | Density (literature): | 1.0275 | mg/µL
| 0,1026 102,90 | 0.1026 | 102.6 | 98.3701 |  | 1,0307 1033,69 | 1.0307 | 1030.7 | 988.2071
| 0,1020 102,30 | 0.102 | 102 | 97.7948 |  | 1,0376 1040,61 | 1.0376 | 1037.6 | 994.8226 | Technique: | Reverse
| 0,1016 101,89 | 0.1016 | 101.6 | 97.4113 |  | 1,0364 1039,41 | 1.0364 | 1036.4 | 993.6721 | Speed: | Fast
| 0,1027 103,00 | 0.1027 | 102.7 | 98.466 |  | 1,0357 1038,70 | 1.0357 | 1035.7 | 993.001
| 0,1028 103,10 | 0.1028 | 102.8 | 98.5618 |  | 1,0340 1037,00 | 1.034 | 1034 | 991.371
| 0,1025 102,80 | 0.1025 | 102.5 | 98.2742 |  | 1,0352 1038,20 | 1.0352 | 1035.2 | 992.5216
| 0,1040 104,30 | 0.104 | 104 | 99.7124 |  | 1,0335 1036,50 | 1.0335 | 1033.5 | 990.8917
| 0,1040 104,30 | 0.104 | 104 | 99.7124 |  | 1,0350 1038,00 | 1.035 | 1035 | 992.3298
| 0,1054 105,71 | 0.1054 | 105.4 | 101.0547 |  | 1,0374 1040,41 | 1.0374 | 1037.4 | 994.6309
| 0,1056 105,91 | 0.1056 | 105.6 | 101.2464 |  | 1,0354 1038,40 | 1.0354 | 1035.4 | 992.7133
| Systematic error |  | µL | -0.9396 |  | Systematic error |  | µL | -7.5839
| Systematic error |  | % | -0.94851 |  | Systematic error |  | % | -0.76418
| Random error |  | µL | 1.3176 |  | Random error |  | µL | 1.9395
| Random error |  | % | 1.33 |  | Random error |  | % | 0.19543
| Uncertainty |  | µL | 3.5747 |  | Uncertainty |  | µL | 11.4629
[5 empty] | 17 | Andrew Alliance Picus 50-1000 µL |  |  | 2016-05-31 | Matrix: | Blood (Sheep, +sitraatti)
[6 empty] | Sartorius Optifit Tip 10-1000 µL |  |  |  |  | Bio Trading BTSN250
[6 empty] | Wide Bore |  | 1000 | 99.88% |  | Lot: #16130 SN78
[6 empty] | Average |  |  | 998.7536 |  | MW:
[6 empty] | Standard deviation |  |  | 2.0722 | Density (experimental): | 1.043 | mg/µL
[7 empty] | [g] | [mg] | [µL] | Density (literature): | 1.0275 | mg/µL
[6 empty] | 1,0385 1041,51 | 1.0385 | 1038.5 | 995.6855
[6 empty] | 1,0401 1043,12 | 1.0401 | 1040.1 | 997.2196 | Technique: | Forward
[6 empty] | 1,0416 1044,62 | 1.0416 | 1041.6 | 998.6577 | Speed: | Slow
[6 empty] | 1,0398 1042,82 | 1.0398 | 1039.8 | 996.9319
[6 empty] | 1,0461 1049,13 | 1.0461 | 1046.1 | 1002.9722
[6 empty] | 1,0429 1045,92 | 1.0429 | 1042.9 | 999.9041
[6 empty] | 1,0416 1044,62 | 1.0416 | 1041.6 | 998.6577
[6 empty] | 1,0435 1046,53 | 1.0435 | 1043.5 | 1000.4794
[6 empty] | 1,0423 1045,32 | 1.0423 | 1042.3 | 999.3289
[6 empty] | 1,0406 1043,62 | 1.0406 | 1040.6 | 997.6989
[6 empty] | Systematic error |  | µL | -1.2464
[6 empty] | Systematic error |  | % | -0.1248
[6 empty] | Random error |  | µL | 2.0722
[6 empty] | Random error |  | % | 0.20748
[6 empty] | Uncertainty |  | µL | 5.3907
[5 empty] | 18 | Andrew Alliance Picus 50-1000 µL |  |  | 2016-05-31 | Matrix: | Blood (Sheep, +sitraatti)
[6 empty] | Sartorius Optifit Tip 10-1000 µL |  |  |  |  | Bio Trading BTSN250
[6 empty] | Wide Bore |  | 1000 | 100.6% |  | Lot: #16130 SN78
[6 empty] | Average |  |  | 1006.4334 |  | MW:
[6 empty] | Standard deviation |  |  | 11.1114 | Density (experimental): | 1.043 | mg/µL
[7 empty] | [g] | [mg] | [µL] | Density (literature): | 1.0275 | mg/µL | < Sheep (vain plasma?)
[6 empty] | 1,0537 1056,76 | 1.0537 | 1053.7 | 1010.2589
[6 empty] | 1,0569 1059,97 | 1.0569 | 1056.9 | 1013.3269 | Technique: | Forward
[6 empty] | 1,0223 1025,26 | 1.0223 | 1022.3 | 980.1534 | Speed: | Fast
[6 empty] | 1,0548 1057,86 | 1.0548 | 1054.8 | 1011.3135
[6 empty] | 1,0553 1058,36 | 1.0553 | 1055.3 | 1011.7929
[6 empty] | 1,0641 1067,19 | 1.0641 | 1064.1 | 1020.2301
[6 empty] | 1,0433 1046,33 | 1.0433 | 1043.3 | 1000.2876
[6 empty] | 1,0416 1044,62 | 1.0416 | 1041.6 | 998.6577
[6 empty] | 1,0514 1054,45 | 1.0514 | 1051.4 | 1008.0537
[6 empty] | 1,0537 1056,76 | 1.0537 | 1053.7 | 1010.2589
[6 empty] | Systematic error |  | µL | 6.4334
[6 empty] | Systematic error |  | % | 0.63922
[6 empty] | Random error |  | µL | 11.1114
[6 empty] | Random error |  | % | 1.104
[6 empty] | Uncertainty |  | µL | 28.6562
[5 empty] | 19 | Andrew Alliance Picus 50-1000 µL |  |  | 2016-05-31 | Matrix: | Blood (Sheep, +sitraatti)
[6 empty] | Sartorius Optifit Tip 10-1000 µL |  |  |  |  | Bio Trading BTSN250
[6 empty] | Wide Bore |  | 1000 | 99.4% |  | Lot: #16130 SN78
[6 empty] | Average |  |  | 993.9789 |  | MW:
[6 empty] | Standard deviation |  |  | 2.5674 | Density (experimental): | 1.043 | mg/µL
[7 empty] | [g] | [mg] | [µL] | Density (literature): | 1.0275 | mg/µL
[6 empty] | 1,0354 1038,40 | 1.0354 | 1035.4 | 992.7133
[6 empty] | 1,0374 1040,41 | 1.0374 | 1037.4 | 994.6309 | Technique: | Reverse
[6 empty] | 1,0392 1042,21 | 1.0392 | 1039.2 | 996.3567 | Speed: | Slow
[6 empty] | 1,0417 1044,72 | 1.0417 | 1041.7 | 998.7536
[6 empty] | 1,0372 1040,21 | 1.0372 | 1037.2 | 994.4391
[6 empty] | 1,0352 1038,20 | 1.0352 | 1035.2 | 992.5216
[6 empty] | 1,0391 1042,11 | 1.0391 | 1039.1 | 996.2608
[6 empty] | 1,0343 1037,30 | 1.0343 | 1034.3 | 991.6587
[6 empty] | 1,0345 1037,50 | 1.0345 | 1034.5 | 991.8504
[6 empty] | 1,0332 1036,20 | 1.0332 | 1033.2 | 990.604
[6 empty] | Systematic error |  | µL | -6.0211
[6 empty] | Systematic error |  | % | -0.60576
[6 empty] | Random error |  | µL | 2.5674
[6 empty] | Random error |  | % | 0.2583
[6 empty] | Uncertainty |  | µL | 11.1559
[5 empty] | 20 | Andrew Alliance Picus 50-1000 µL |  |  |  | Matrix: | Blood (Sheep, +sitraatti)
[6 empty] | Sartorius Optifit Tip 10-1000 µL |  |  |  |  | Bio Trading BTSN250
[6 empty] | Wide Bore |  | 1000 | 99.08% |  | Lot: #16130 SN78
[6 empty] | Average |  |  | 990.8054 |  | MW:
[6 empty] | Standard deviation |  |  | 2.904 | Density (experimental): | 1.043 | mg/µL
[7 empty] | [g] | [mg] | [µL] | Density (literature): | 1.0275 | mg/µL | < Sheep (vain plasma?)
[6 empty] | 1,0329 1035,90 | 1.0329 | 1032.9 | 990.3164
[6 empty] | 1,0289 1031,88 | 1.0289 | 1028.9 | 986.4813 | Technique: | Reverse
[6 empty] | 1,0326 1035,59 | 1.0326 | 1032.6 | 990.0288 | Speed: | Fast
[6 empty] | 1,0306 1033,59 | 1.0306 | 1030.6 | 988.1112
[6 empty] | 1,0371 1040,11 | 1.0371 | 1037.1 | 994.3432
[6 empty] | 1,0311 1034,09 | 1.0311 | 1031.1 | 988.5906
[6 empty] | 1,0357 1038,70 | 1.0357 | 1035.7 | 993.001
[6 empty] | 1,0330 1036,00 | 1.033 | 1033 | 990.4123
[6 empty] | 1,0334 1036,40 | 1.0334 | 1033.4 | 990.7958
[6 empty] | 1,0388 1041,81 | 1.0388 | 1038.8 | 995.9732
[6 empty] | Systematic error |  | µL | -9.1946
[6 empty] | Systematic error |  | % | -0.928
[6 empty] | Random error |  | µL | 2.904
[6 empty] | Random error |  | % | 0.2931
[6 empty] | Uncertainty |  | µL | 15.0027

## Methanol
1 | Andrew Alliance Picus 5-120 µL |  |  |  | 9 | Andrew Alliance Picus 50-1000 µL |  |  |  | Matrix: | Methanol
| Sartorius Optifit Tip 0,5-200 µL |  |  |  |  | Sartorius Optifit Tip 10-1000 µL |  |  |  |  | Fisher Chemical M/4056
|  |  | 100 | 97.71% |  |  |  | 1000 | 100.2% |  | Lot #1668488
| Average |  |  | 97.7089 |  | Average |  |  | 1002.0886 |  | MW: 32,04
| Standard deviation |  |  | 0.54183 |  | Standard deviation |  |  | 3.1458 | Density (experimental): | 0.79 | mg/µL
|  | [g] | [mg] | [µL] |  |  | [g] | [mg] | [µL] | Density (literature): | 0.791 | mg/µL
| 0,0779 78,13 | 0.0779 | 77.9 | 98.6076 |  | 0,7933 795,60 | 0.7933 | 793.3 | 1004.1772
| 0,0773 77,52 | 0.0773 | 77.3 | 97.8481 |  | 0,7963 798,61 | 0.7963 | 796.3 | 1007.9747 | Technique: | Forward
| 0,0778 78,03 | 0.0778 | 77.8 | 98.481 |  | 0,7911 793,39 | 0.7911 | 791.1 | 1001.3924 | Speed: | Slow
| 0,0769 77,12 | 0.0769 | 76.9 | 97.3418 |  | 0,7926 794,90 | 0.7926 | 792.6 | 1003.2911
| 0,0771 77,32 | 0.0771 | 77.1 | 97.5949 |  | 0,7918 794,10 | 0.7918 | 791.8 | 1002.2785
| 0,0764 76,62 | 0.0764 | 76.4 | 96.7089 |  | 0,7929 795,20 | 0.7929 | 792.9 | 1003.6709
| 0,0770 77,22 | 0.077 | 77 | 97.4684 |  | 0,7908 793,09 | 0.7908 | 790.8 | 1001.0127
| 0,0772 77,42 | 0.0772 | 77.2 | 97.7215 |  | 0,7871 789,38 | 0.7871 | 787.1 | 996.3291
| 0,0772 77,42 | 0.0772 | 77.2 | 97.7215 |  | 0,7890 791,29 | 0.789 | 789 | 998.7342
| 0,0771 77,32 | 0.0771 | 77.1 | 97.5949 |  | 0,7916 793,90 | 0.7916 | 791.6 | 1002.0253
| Systematic error |  | µL | -2.2911 |  | Systematic error |  | µL | 2.0886
| Systematic error |  | % | -2.3449 |  | Systematic error |  | % | 0.20843
| Random error |  | µL | 0.54183 |  | Random error |  | µL | 3.1458
| Random error |  | % | 0.55453 |  | Random error |  | % | 0.31392
| Uncertainty |  | µL | 3.3748 |  | Uncertainty |  | µL | 8.3802
2 | Andrew Alliance Picus 5-120 µL |  |  |  | 10 | Andrew Alliance Picus 50-1000 µL |  |  |  | Matrix: | Methanol
| Sartorius Optifit Tip 0,5-200 µL |  |  |  |  | Sartorius Optifit Tip 10-1000 µL |  |  |  |  | Fisher Chemical M/4056
|  |  | 100 | 96.65% |  |  |  | 1000 | 99.35% |  | Lot #1668488
| Average |  |  | 96.6456 |  | Average |  |  | 993.481 |  | MW: 32,04
| Standard deviation |  |  | 0.36419 |  | Standard deviation |  |  | 3.2285 | Density (experimental): | 0.79 | mg/µL
|  | [g] | [mg] | [µL] |  |  | [g] | [mg] | [µL] | Density (literature): | 0.791 | mg/µL
| 0,0764 76,62 | 0.0764 | 76.4 | 96.7089 |  | 0,7843 786,57 | 0.7843 | 784.3 | 992.7848
| 0,0765 76,72 | 0.0765 | 76.5 | 96.8354 |  | 0,7886 790,89 | 0.7886 | 788.6 | 998.2278 | Technique: | Forward
| 0,0765 76,72 | 0.0765 | 76.5 | 96.8354 |  | 0,7795 781,76 | 0.7795 | 779.5 | 986.7089 | Speed: | Fast
| 0,0768 77,02 | 0.0768 | 76.8 | 97.2152 |  | 0,7854 787,68 | 0.7854 | 785.4 | 994.1772
| 0,0762 76,42 | 0.0762 | 76.2 | 96.4557 |  | 0,7854 787,68 | 0.7854 | 785.4 | 994.1772
| 0,0764 76,62 | 0.0764 | 76.4 | 96.7089 |  | 0,7833 785,57 | 0.7833 | 783.3 | 991.519
| 0,0759 76,12 | 0.0759 | 75.9 | 96.0759 |  | 0,7878 790,08 | 0.7878 | 787.8 | 997.2152
| 0,0766 76,82 | 0.0766 | 76.6 | 96.962 |  | 0,7842 786,47 | 0.7842 | 784.2 | 992.6582
| 0,0759 76,12 | 0.0759 | 75.9 | 96.0759 |  | 0,7837 785,97 | 0.7837 | 783.7 | 992.0253
| 0,0763 76,52 | 0.0763 | 76.3 | 96.5823 |  | 0,7863 788,58 | 0.7863 | 786.3 | 995.3165
| Systematic error |  | µL | -3.3544 |  | Systematic error |  | µL | -6.519
| Systematic error |  | % | -3.4709 |  | Systematic error |  | % | -0.65618
| Random error |  | µL | 0.36419 |  | Random error |  | µL | 3.2285
| Random error |  | % | 0.37683 |  | Random error |  | % | 0.32497
| Uncertainty |  | µL | 4.0828 |  | Uncertainty |  | µL | 12.9759
3 | Andrew Alliance Picus 5-120 µL |  |  |  | 11 | Andrew Alliance Picus 50-1000 µL |  |  |  | Matrix: | Methanol
| Sartorius Optifit Tip 0,5-200 µL |  |  |  |  | Sartorius Optifit Tip 10-1000 µL |  |  |  |  | Fisher Chemical M/4056
|  |  | 100 | 98.24% |  |  |  | 1000 | 99.38% |  | Lot #1668488
| Average |  |  | 98.2405 |  | Average |  |  | 993.8228 |  | MW: 32,04
| Standard deviation |  |  | 0.4562 |  | Standard deviation |  |  | 1.8592 | Density (experimental): | 0.79 | mg/µL
|  | [g] | [mg] | [µL] |  |  | [g] | [mg] | [µL] | Density (literature): | 0.791 | mg/µL
| 0,0775 77,72 | 0.0775 | 77.5 | 98.1013 |  | 0,7848 787,08 | 0.7848 | 784.8 | 993.4177
| 0,0784 78,63 | 0.0784 | 78.4 | 99.2405 |  | 0,7857 787,98 | 0.7857 | 785.7 | 994.557 | Technique: | Reverse
| 0,0779 78,13 | 0.0779 | 77.9 | 98.6076 |  | 0,7847 786,98 | 0.7847 | 784.7 | 993.2911 | Speed: | Slow
| 0,0777 77,93 | 0.0777 | 77.7 | 98.3544 |  | 0,7861 788,38 | 0.7861 | 786.1 | 995.0633
| 0,0776 77,83 | 0.0776 | 77.6 | 98.2278 |  | 0,7852 787,48 | 0.7852 | 785.2 | 993.9241
| 0,0770 77,22 | 0.077 | 77 | 97.4684 |  | 0,7860 788,28 | 0.786 | 786 | 994.9367
| 0,0776 77,83 | 0.0776 | 77.6 | 98.2278 |  | 0,7819 784,17 | 0.7819 | 781.9 | 989.7468
| 0,0775 77,72 | 0.0775 | 77.5 | 98.1013 |  | 0,7836 785,87 | 0.7836 | 783.6 | 991.8987
| 0,0774 77,62 | 0.0774 | 77.4 | 97.9747 |  | 0,7866 788,88 | 0.7866 | 786.6 | 995.6962
| 0,0775 77,72 | 0.0775 | 77.5 | 98.1013 |  | 0,7866 788,88 | 0.7866 | 786.6 | 995.6962
| Systematic error |  | µL | -1.7595 |  | Systematic error |  | µL | -6.1772
| Systematic error |  | % | -1.791 |  | Systematic error |  | % | -0.62156
| Random error |  | µL | 0.4562 |  | Random error |  | µL | 1.8592
| Random error |  | % | 0.46437 |  | Random error |  | % | 0.18708
| Uncertainty |  | µL | 2.6719 |  | Uncertainty |  | µL | 9.8957
4 | Andrew Alliance Picus 5-120 µL |  |  |  | 12 | Andrew Alliance Picus 50-1000 µL |  |  |  | Matrix: | Methanol
| Sartorius Optifit Tip 0,5-200 µL |  |  |  |  | Sartorius Optifit Tip 10-1000 µL |  |  |  |  | Fisher Chemical M/4056
|  |  | 100 | 100.1% |  |  |  | 1000 | 98.36% |  | Lot #1668488
| Average |  |  | 100.0886 |  | Average |  |  | 983.6329 |  | MW: 32,04
| Standard deviation |  |  | 0.43462 |  | Standard deviation |  |  | 1.3169 | Density (experimental): | 0.79 | mg/µL
|  | [g] | [mg] | [µL] |  |  | [g] | [mg] | [µL] | Density (literature): | 0.791 | mg/µL
| 0,0784 78,63 | 0.0784 | 78.4 | 99.2405 |  | 0,7773 779,55 | 0.7773 | 777.3 | 983.9241
| 0,0791 79,33 | 0.0791 | 79.1 | 100.1266 |  | 0,7775 779,75 | 0.7775 | 777.5 | 984.1772 | Technique: | Reverse
| 0,0794 79,63 | 0.0794 | 79.4 | 100.5063 |  | 0,7772 779,45 | 0.7772 | 777.2 | 983.7975 | Speed: | Fast
| 0,0787 78,93 | 0.0787 | 78.7 | 99.6203 |  | 0,7765 778,75 | 0.7765 | 776.5 | 982.9114
| 0,0792 79,43 | 0.0792 | 79.2 | 100.2532 |  | 0,7775 779,75 | 0.7775 | 777.5 | 984.1772
| 0,0796 79,83 | 0.0796 | 79.6 | 100.7595 |  | 0,7782 780,46 | 0.7782 | 778.2 | 985.0633
| 0,0792 79,43 | 0.0792 | 79.2 | 100.2532 |  | 0,7781 780,36 | 0.7781 | 778.1 | 984.9367
| 0,0792 79,43 | 0.0792 | 79.2 | 100.2532 |  | 0,7770 779,25 | 0.777 | 777 | 983.5443
| 0,0789 79,13 | 0.0789 | 78.9 | 99.8734 |  | 0,7745 776,75 | 0.7745 | 774.5 | 980.3797
| 0,0790 79,23 | 0.079 | 79 | 100 |  | 0,7769 779,15 | 0.7769 | 776.9 | 983.4177
| Systematic error |  | µL | 0.088608 |  | Systematic error |  | µL | -16.3671
| Systematic error |  | % | 0.088529 |  | Systematic error |  | % | -1.6639
| Random error |  | µL | 0.43462 |  | Random error |  | µL | 1.3169
| Random error |  | % | 0.43423 |  | Random error |  | % | 0.13388
| Uncertainty |  | µL | 0.95785 |  | Uncertainty |  | µL | 19.0009
5 | Andrew Alliance Picus 5-120 µL |  |  |  | 13 | Andrew Alliance Picus 50-1000 µL |  |  |  | Matrix: | Methanol
| Sartorius Optifit Tip 0,5-200 µL |  |  |  |  | Sartorius Optifit Tip 10-1000 µL |  |  |  |  | Fisher Chemical M/4056
| Low Retention |  | 100 | 97.28% |  | Low Retention |  | 1000 | 99.55% |  | Lot #1668488
| Average |  |  | 97.2785 |  | Average |  |  | 995.5443 |  | MW: 32,04
| Standard deviation |  |  | 0.61508 |  | Standard deviation |  |  | 2.6957 | Density (experimental): | 0.79 | mg/µL
|  | [g] | [mg] | [µL] |  |  | [g] | [mg] | [µL] | Density (literature): | 0.791 | mg/µL
| 0,0757 75,92 | 0.0757 | 75.7 | 95.8228 |  | 0,7817 783,97 | 0.7817 | 781.7 | 989.4937
| 0,0768 77,02 | 0.0768 | 76.8 | 97.2152 |  | 0,7844 786,67 | 0.7844 | 784.4 | 992.9114 | Technique: | Forward
| 0,0765 76,72 | 0.0765 | 76.5 | 96.8354 |  | 0,7852 787,48 | 0.7852 | 785.2 | 993.9241 | Speed: | Slow
| 0,0771 77,32 | 0.0771 | 77.1 | 97.5949 |  | 0,7880 790,29 | 0.788 | 788 | 997.4684
| 0,0769 77,12 | 0.0769 | 76.9 | 97.3418 |  | 0,7878 790,08 | 0.7878 | 787.8 | 997.2152
| 0,0772 77,42 | 0.0772 | 77.2 | 97.7215 |  | 0,7871 789,38 | 0.7871 | 787.1 | 996.3291
| 0,0771 77,32 | 0.0771 | 77.1 | 97.5949 |  | 0,7868 789,08 | 0.7868 | 786.8 | 995.9494
| 0,0769 77,12 | 0.0769 | 76.9 | 97.3418 |  | 0,7879 790,18 | 0.7879 | 787.9 | 997.3418
| 0,0768 77,02 | 0.0768 | 76.8 | 97.2152 |  | 0,7872 789,48 | 0.7872 | 787.2 | 996.4557
| 0,0775 77,72 | 0.0775 | 77.5 | 98.1013 |  | 0,7887 790,99 | 0.7887 | 788.7 | 998.3544
| Systematic error |  | µL | -2.7215 |  | Systematic error |  | µL | -4.4557
| Systematic error |  | % | -2.7977 |  | Systematic error |  | % | -0.44756
| Random error |  | µL | 0.61508 |  | Random error |  | µL | 2.6957
| Random error |  | % | 0.63229 |  | Random error |  | % | 0.27077
| Uncertainty |  | µL | 3.9517 |  | Uncertainty |  | µL | 9.847
6 | Andrew Alliance Picus 5-120 µL |  |  |  | 14 | Andrew Alliance Picus 50-1000 µL |  |  |  | Matrix: | Methanol
| Sartorius Optifit Tip 0,5-200 µL |  |  |  |  | Sartorius Optifit Tip 10-1000 µL |  |  |  |  | Fisher Chemical M/4056
| Low Retention |  | 100 | 99.71% |  | Low Retention |  | 1000 | 99.62% |  | Lot #1668488
| Average |  |  | 99.7089 |  | Average |  |  | 996.2278 |  | MW: 32,04
| Standard deviation |  |  | 0.49585 |  | Standard deviation |  |  | 2.0548 | Density (experimental): | 0.79 | mg/µL
|  | [g] | [mg] | [µL] |  |  | [g] | [mg] | [µL] | Density (literature): | 0.791 | mg/µL
| 0,0788 79,03 | 0.0788 | 78.8 | 99.7468 |  | 0,7879 790,18 | 0.7879 | 787.9 | 997.3418
| 0,0789 79,13 | 0.0789 | 78.9 | 99.8734 |  | 0,7886 790,89 | 0.7886 | 788.6 | 998.2278 | Technique: | Forward
| 0,0789 79,13 | 0.0789 | 78.9 | 99.8734 |  | 0,7853 787,58 | 0.7853 | 785.3 | 994.0506 | Speed: | Fast
| 0,0791 79,33 | 0.0791 | 79.1 | 100.1266 |  | 0,7882 790,49 | 0.7882 | 788.2 | 997.7215
| 0,0791 79,33 | 0.0791 | 79.1 | 100.1266 |  | 0,7842 786,47 | 0.7842 | 784.2 | 992.6582
| 0,0781 78,33 | 0.0781 | 78.1 | 98.8608 |  | 0,7871 789,38 | 0.7871 | 787.1 | 996.3291
| 0,0789 79,13 | 0.0789 | 78.9 | 99.8734 |  | 0,7880 790,29 | 0.788 | 788 | 997.4684
| 0,0780 78,23 | 0.078 | 78 | 98.7342 |  | 0,7878 790,08 | 0.7878 | 787.8 | 997.2152
| 0,0790 79,23 | 0.079 | 79 | 100 |  | 0,7848 787,08 | 0.7848 | 784.8 | 993.4177
| 0,0789 79,13 | 0.0789 | 78.9 | 99.8734 |  | 0,7883 790,59 | 0.7883 | 788.3 | 997.8481
| Systematic error |  | µL | -0.29114 |  | Systematic error |  | µL | -3.7722
| Systematic error |  | % | -0.29199 |  | Systematic error |  | % | -0.37864
| Random error |  | µL | 0.49585 |  | Random error |  | µL | 2.0548
| Random error |  | % | 0.4973 |  | Random error |  | % | 0.20626
| Uncertainty |  | µL | 1.2828 |  | Uncertainty |  | µL | 7.8818
7 | Andrew Alliance Picus 5-120 µL |  |  |  | 15 | Andrew Alliance Picus 50-1000 µL |  |  |  | Matrix: | Methanol
| Sartorius Optifit Tip 0,5-200 µL |  |  |  |  | Sartorius Optifit Tip 10-1000 µL |  |  |  |  | Fisher Chemical M/4056
| Low Retention |  | 100 | 97.89% |  | Low Retention |  | 1000 | 99.53% |  | Lot #1668488
| Average |  |  | 97.8861 |  | Average |  |  | 995.3418 |  | MW: 32,04
| Standard deviation |  |  | 0.63164 |  | Standard deviation |  |  | 0.8307 | Density (experimental): | 0.79 | mg/µL
|  | [g] | [mg] | [µL] |  |  | [g] | [mg] | [µL] | Density (literature): | 0.791 | mg/µL
| 0,0770 77,22 | 0.077 | 77 | 97.4684 |  | 0,7868 789,08 | 0.7868 | 786.8 | 995.9494
| 0,0779 78,13 | 0.0779 | 77.9 | 98.6076 |  | 0,7875 789,78 | 0.7875 | 787.5 | 996.8354 | Technique: | Reverse
| 0,0767 76,92 | 0.0767 | 76.7 | 97.0886 |  | 0,7867 788,98 | 0.7867 | 786.7 | 995.8228 | Speed: | Slow
| 0,0768 77,02 | 0.0768 | 76.8 | 97.2152 |  | 0,7859 788,18 | 0.7859 | 785.9 | 994.8101
| 0,0767 76,92 | 0.0767 | 76.7 | 97.0886 |  | 0,7856 787,88 | 0.7856 | 785.6 | 994.4304
| 0,0776 77,83 | 0.0776 | 77.6 | 98.2278 |  | 0,7871 789,38 | 0.7871 | 787.1 | 996.3291
| 0,0780 78,23 | 0.078 | 78 | 98.7342 |  | 0,7861 788,38 | 0.7861 | 786.1 | 995.0633
| 0,0777 77,93 | 0.0777 | 77.7 | 98.3544 |  | 0,7860 788,28 | 0.786 | 786 | 994.9367
| 0,0776 77,83 | 0.0776 | 77.6 | 98.2278 |  | 0,7858 788,08 | 0.7858 | 785.8 | 994.6835
| 0,0773 77,52 | 0.0773 | 77.3 | 97.8481 |  | 0,7857 787,98 | 0.7857 | 785.7 | 994.557
| Systematic error |  | µL | -2.1139 |  | Systematic error |  | µL | -4.6582
| Systematic error |  | % | -2.1596 |  | Systematic error |  | % | -0.468
| Random error |  | µL | 0.63164 |  | Random error |  | µL | 0.8307
| Random error |  | % | 0.64529 |  | Random error |  | % | 0.083459
| Uncertainty |  | µL | 3.3772 |  | Uncertainty |  | µL | 6.3196
8 | Andrew Alliance Picus 5-120 µL |  |  |  | 16 | Andrew Alliance Picus 50-1000 µL |  |  |  | Matrix: | Methanol
| Sartorius Optifit Tip 0,5-200 µL |  |  |  |  | Sartorius Optifit Tip 10-1000 µL |  |  |  |  | Fisher Chemical M/4056
| Low Retention |  | 100 | 98.47% |  | Low Retention |  | 1000 | 99.33% |  | Lot #1668488
| Average |  |  | 98.4684 |  | Average |  |  | 993.2785 |  | MW: 32,04
| Standard deviation |  |  | 0.5516 |  | Standard deviation |  |  | 0.52514 | Density (experimental): | 0.79 | mg/µL
|  | [g] | [mg] | [µL] |  |  | [g] | [mg] | [µL] | Density (literature): | 0.791 | mg/µL
| 0,0774 77,62 | 0.0774 | 77.4 | 97.9747 |  | 0,7845 786,78 | 0.7845 | 784.5 | 993.038
| 0,0776 77,83 | 0.0776 | 77.6 | 98.2278 |  | 0,7846 786,88 | 0.7846 | 784.6 | 993.1646 | Technique: | Reverse
| 0,0783 78,53 | 0.0783 | 78.3 | 99.1139 |  | 0,7849 787,18 | 0.7849 | 784.9 | 993.5443 | Speed: | Fast
| 0,0772 77,42 | 0.0772 | 77.2 | 97.7215 |  | 0,7855 787,78 | 0.7855 | 785.5 | 994.3038
| 0,0781 78,33 | 0.0781 | 78.1 | 98.8608 |  | 0,7851 787,38 | 0.7851 | 785.1 | 993.7975
| 0,0778 78,03 | 0.0778 | 77.8 | 98.481 |  | 0,7844 786,67 | 0.7844 | 784.4 | 992.9114
| 0,0782 78,43 | 0.0782 | 78.2 | 98.9873 |  | 0,7850 787,28 | 0.785 | 785 | 993.6709
| 0,0773 77,52 | 0.0773 | 77.3 | 97.8481 |  | 0,7843 786,57 | 0.7843 | 784.3 | 992.7848
| 0,0784 78,63 | 0.0784 | 78.4 | 99.2405 |  | 0,7843 786,57 | 0.7843 | 784.3 | 992.7848
| 0,0776 77,83 | 0.0776 | 77.6 | 98.2278 |  | 0,7843 786,57 | 0.7843 | 784.3 | 992.7848
| Systematic error |  | µL | -1.5316 |  | Systematic error |  | µL | -6.7215
| Systematic error |  | % | -1.5555 |  | Systematic error |  | % | -0.6767
| Random error |  | µL | 0.5516 |  | Random error |  | µL | 0.52514
| Random error |  | % | 0.56018 |  | Random error |  | % | 0.05287
| Uncertainty |  | µL | 2.6348 |  | Uncertainty |  | µL | 7.7718
[5 empty] | 17 | Andrew Alliance Picus 50-1000 µL |  |  |  | Matrix: | Methanol
[6 empty] | Sartorius Optifit Tip 10-1000 µL |  |  |  |  | Fisher Chemical M/4056
[6 empty] | Wide Bore |  | 1000 | 100.2% |  | Lot #1668488
[6 empty] | Average |  |  | 1001.7975 |  | MW: 32,04
[6 empty] | Standard deviation |  |  | 1.8592 | Density (experimental): | 0.79 | mg/µL
[7 empty] | [g] | [mg] | [µL] | Density (literature): | 0.791 | mg/µL
[6 empty] | 0,7924 794,70 | 0.7924 | 792.4 | 1003.038
[6 empty] | 0,7938 796,10 | 0.7938 | 793.8 | 1004.8101 | Technique: | Forward
[6 empty] | 0,7937 796,00 | 0.7937 | 793.7 | 1004.6835 | Speed: | Slow
[6 empty] | 0,7907 792,99 | 0.7907 | 790.7 | 1000.8861
[6 empty] | 0,7900 792,29 | 0.79 | 790 | 1000
[6 empty] | 0,7911 793,39 | 0.7911 | 791.1 | 1001.3924
[6 empty] | 0,7919 794,20 | 0.7919 | 791.9 | 1002.4051
[6 empty] | 0,7899 792,19 | 0.7899 | 789.9 | 999.8734
[6 empty] | 0,7901 792,39 | 0.7901 | 790.1 | 1000.1266
[6 empty] | 0,7906 792,89 | 0.7906 | 790.6 | 1000.7595
[6 empty] | Systematic error |  | µL | 1.7975
[6 empty] | Systematic error |  | % | 0.17942
[6 empty] | Random error |  | µL | 1.8592
[6 empty] | Random error |  | % | 0.18559
[6 empty] | Uncertainty |  | µL | 5.5159
[5 empty] | 18 | Andrew Alliance Picus 50-1000 µL |  |  |  | Matrix: | Methanol
[6 empty] | Sartorius Optifit Tip 10-1000 µL |  |  |  |  | Fisher Chemical M/4056
[6 empty] | Wide Bore |  | 1000 | 98.81% |  | Lot #1668488
[6 empty] | Average |  |  | 988.1139 |  | MW: 32,04
[6 empty] | Standard deviation |  |  | 5.015 | Density (experimental): | 0.79 | mg/µL
[7 empty] | [g] | [mg] | [µL] | Density (literature): | 0.791 | mg/µL
[6 empty] | 0,7855 787,78 | 0.7855 | 785.5 | 994.3038
[6 empty] | 0,7772 779,45 | 0.7772 | 777.2 | 983.7975 | Technique: | Forward
[6 empty] | 0,7788 781,06 | 0.7788 | 778.8 | 985.8228 | Speed: | Fast
[6 empty] | 0,7803 782,56 | 0.7803 | 780.3 | 987.7215
[6 empty] | 0,7782 780,46 | 0.7782 | 778.2 | 985.0633
[6 empty] | 0,7792 781,46 | 0.7792 | 779.2 | 986.3291
[6 empty] | 0,7792 781,46 | 0.7792 | 779.2 | 986.3291
[6 empty] | 0,7768 779,05 | 0.7768 | 776.8 | 983.2911
[6 empty] | 0,7815 783,77 | 0.7815 | 781.5 | 989.2405
[6 empty] | 0,7894 791,69 | 0.7894 | 789.4 | 999.2405
[6 empty] | Systematic error |  | µL | -11.8861
[6 empty] | Systematic error |  | % | -1.2029
[6 empty] | Random error |  | µL | 5.015
[6 empty] | Random error |  | % | 0.50754
[6 empty] | Uncertainty |  | µL | 21.9162
[5 empty] | 19 | Andrew Alliance Picus 50-1000 µL |  |  |  | Matrix: | Methanol
[6 empty] | Sartorius Optifit Tip 10-1000 µL |  |  |  |  | Fisher Chemical M/4056
[6 empty] | Wide Bore |  | 1000 | 99.32% |  | Lot #1668488
[6 empty] | Average |  |  | 993.2405 |  | MW: 32,04
[6 empty] | Standard deviation |  |  | 1.1246 | Density (experimental): | 0.79 | mg/µL
[7 empty] | [g] | [mg] | [µL] | Density (literature): | 0.791 | mg/µL
[6 empty] | 0,7859 788,18 | 0.7859 | 785.9 | 994.8101
[6 empty] | 0,7854 787,68 | 0.7854 | 785.4 | 994.1772 | Technique: | Reverse
[6 empty] | 0,7857 787,98 | 0.7857 | 785.7 | 994.557 | Speed: | Slow
[6 empty] | 0,7840 786,27 | 0.784 | 784 | 992.4051
[6 empty] | 0,7830 785,27 | 0.783 | 783 | 991.1392
[6 empty] | 0,7841 786,37 | 0.7841 | 784.1 | 992.5316
[6 empty] | 0,7849 787,18 | 0.7849 | 784.9 | 993.5443
[6 empty] | 0,7841 786,37 | 0.7841 | 784.1 | 992.5316
[6 empty] | 0,7846 786,88 | 0.7846 | 784.6 | 993.1646
[6 empty] | 0,7849 787,18 | 0.7849 | 784.9 | 993.5443
[6 empty] | Systematic error |  | µL | -6.7595
[6 empty] | Systematic error |  | % | -0.68055
[6 empty] | Random error |  | µL | 1.1246
[6 empty] | Random error |  | % | 0.11323
[6 empty] | Uncertainty |  | µL | 9.0087
[5 empty] | 20 | Andrew Alliance Picus 50-1000 µL |  |  |  | Matrix: | Methanol
[6 empty] | Sartorius Optifit Tip 10-1000 µL |  |  |  |  | Fisher Chemical M/4056
[6 empty] | Wide Bore |  | 1000 | 98.78% |  | Lot #1668488
[6 empty] | Average |  |  | 987.8354 |  | MW: 32,04
[6 empty] | Standard deviation |  |  | 2.0397 | Density (experimental): | 0.79 | mg/µL
[7 empty] | [g] | [mg] | [µL] | Density (literature): | 0.791 | mg/µL
[6 empty] | 0,7835 785,77 | 0.7835 | 783.5 | 991.7722
[6 empty] | 0,7831 785,37 | 0.7831 | 783.1 | 991.2658 | Technique: | Reverse
[6 empty] | 0,7800 782,26 | 0.78 | 780 | 987.3418 | Speed: | Fast
[6 empty] | 0,7805 782,76 | 0.7805 | 780.5 | 987.9747
[6 empty] | 0,7795 781,76 | 0.7795 | 779.5 | 986.7089
[6 empty] | 0,7801 782,36 | 0.7801 | 780.1 | 987.4684
[6 empty] | 0,7799 782,16 | 0.7799 | 779.9 | 987.2152
[6 empty] | 0,7793 781,56 | 0.7793 | 779.3 | 986.4557
[6 empty] | 0,7789 781,16 | 0.7789 | 778.9 | 985.9494
[6 empty] | 0,7791 781,36 | 0.7791 | 779.1 | 986.2025
[6 empty] | Systematic error |  | µL | -12.1646
[6 empty] | Systematic error |  | % | -1.2314
[6 empty] | Random error |  | µL | 2.0397
[6 empty] | Random error |  | % | 0.20648
[6 empty] | Uncertainty |  | µL | 16.244